## Slide 1: Research Progress on Infrared Benchmark Payload： Highly accurate FTIR
Lei DING1,3, Shijie XIN1,3, Minjian GU1，Xiaopeng HAO2 ，etc.
1 Shanghai Institute of Technical Physics, CAS
2 National Institute of Metrology
3 National Key Laboratory of Infrared Physics
March 2025

## Slide 2: Infrared sounder for FY-3
IRAS
Spectral range: 0.69 ~ 15.5 μm, channel numbers: 26,
cross-track scanning: ±49.5°
(2172 km),
Spatial resolution: 17.0 km
 HIRAS
channel numbers: >1300
Spectral resolution: 0.625cm-1
Spatial resolution: 17.0 km
Accruacy: <1K
2

## Slide 3: Infrared sounder for FY-3
Band | Spectral Range (cm-1) | Spectral Resolution (cm-1) | Sensitivity (NET@280K) |  |  | Num of Channels
 |  |  | FY-3D | FY-3E |  | 
LWIR | 650～1136 （15.38m～8.8 m） | 0.625 | 0.15(Expectation) 0.4K(Requirement) | 650 ～667 cm-1 | 0.8K | 778
 |  |  |  | 667～689 cm-1 | 0.4K | 
 |  |  |  | 689～1000 cm-1 | 0.2K | 
 |  |  |  | 1000～1136 cm-1 | 0.4K | 
MWIR1 | 1210～1750 （8.26m～5.71 m） | 1.25 | 0.1(Expectation) 0.7K(Requirement) | 1210～1538 cm-1 | 0.2K | 433
 |  |  |  | 1538～1750 cm-1 | 0.3K | 
MWIR2 | 2155～2550 （4.64m～3.92 m） | 2.5 | 0.3(Expectation) 1.2K(Requirement) | 2155～2300 cm-1 | 0.3K | 159
 |  |  |  | 2300～2550 cm-1 | 0.5K | 
HIRAS instrument specification improvement for FY-3 D/E/F
3

## Slide 4: Infrared sounder for FY-4
Geostationary Interfering InfraRed Sounder (GIIRS)
Recording
4.44μm－6.06μm、8.85μm－14.3μm
Resolution: 0.625 cm-1
Spectral stability：<10ppm
67min, 5000km×5000km
Calibration accuracy 0.7K@300K
4

## Slide 5: Steps for CRABS
From P. Zhang 2019, at NPL
Development of Chinese RAdiometric Benchmark Satellite
CRABS
From P. Zhang
5
Infrared Spectral Segment

## Slide 6: Space Radiometric Benchmark Reference of Infrared Spectral Segment
Space-borne Infrared Radometric Reference Payload
Infrared Radiometric Benchmark Reference Source
Benchmark Reference Satellite
On-orbit calibration
Benchmark transmission
On-orbit IR source
SI trace in laboratory
Traceable FTIR Prototype
On orbit payload
Remote sensing instruments must be accurate and stable - quantification of remote sensing data
6

## Slide 7: Space Radiometric Benchmark Reference of Infrared Spectral Segment
FTIR
Traceable
Source
Calibration sources
Calibration
Model
RAW
DATA
SI Traceable DATA
sensors
tobe
calibrated
Inter-calibration
SI基准星地传递
Traceable Infrared Spectral Segment
辐射定标基准传递
On-orbit
Application
Flight model for reference satellite
Groud
Facility
7
Notes: 第三部分是总体的研究路线
利用中国计量院的地面基准装置，实现SI基准的星地间传递，红外光谱仪获取原始的观测数据，通过仪器定标和误差订正模型，得到SI可溯源的定标数据，再利用交叉定标及辐射基准传递模型，最终实现红外辐射测量基准载荷与待测仪器间的辐射定标基准传递。

## Slide 8: Space Radiometric Benchmark Reference of Infrared Spectral Segment
Key technologies for Infrared SI-traceable payload
Hyper-spectral infrared sounder
  Interferometer for space
  Detectors
  Temperature control
SI traceable blackbody
  miniature fixed point
  SI benchmark transfer
Vacuum Radiance Temperature Standard Facility
8

## Slide 9: Space Radiometric Benchmark Reference of Infrared Spectral Segment
Several modes
9
Technical Approach
Notes: 左图是载荷电子学设计，其中包括二维扫描镜控制、干涉仪动镜控制、制冷系统控制、干涉信号获取及电子学管理与检测等。
右侧是工作模式，分为对地观测模式、自校准模式和溯源模式三种
这张是载荷结构示意图，分为空间红外辐射基准源、空间红外辐射变温定标源、低温定标源、望远镜、冷箱、制冷机、扫描镜和干涉仪

## Slide 10
Uncertainty 0.085K (k=1)
Foreward interferometer
0.5cm-1 resolution (±10mm)
0.852 μm diode laser for interferogram sample control
< 1ppm
10
Space Radiometric Benchmark Reference of Infrared Spectral Segment

## Slide 11: Space Radiometric Benchmark Reference of Infrared Spectral Segment
Scientific goals
Multi-source infrared remote sensing data is difficult to trace back to a common spectral radiance benchmark
Key techno-logies
Research content
Spurious Immuno-Optical Surface Materials
Temperature correlation control of multi-temperature zone structures
Lightweight integration of traceability structure
Flutter Isolation Structure
Rapid traceability process design
Mid- and long-wave infrared high-sensitivity detector
Nonlinear compensation
Low temperature integrated low power preamplifier
Develop an infrared spectral radiation benchmark payload engineering prototype that can work in a vacuum and low temperature environment
Uncertainty evaluation model of infrared radiation source traceability based on cross-calibration time and space parameters
Weak signal detection
Interfero-meter
Optical system
Sensor
Electronics
Mechanics
Thermal Control
Radiation reference source
Technical Approach

## Slide 12: Space Radiometric Benchmark Reference of Infrared Spectral Segment
IR detector
空间复杂外部杂散光识别与抑制
光学设计和仿真
Optics
甚长波红外
探测器
响应非线性物理机制和测试
低温前置放大和集成
Mechanics
震动耦合链路仿真与抑制
结构设计和仿真
Electronics
Interferometer
Temperature control
微弱信号极限探测原理
全年外热流仿真
多温区热控仿真
12
Technical Approach
Notes: 对载荷关键子部件的技术难点识别与攻关，最终可实现红外辐射基准载荷的红外探测器、分束器、光学薄膜等全部组件的自主研制

## Slide 13: Space Radiometric Benchmark Reference of Infrared Spectral Segment
Uncertainty evaluation of space infrared radiation benchmark payload
Uncertainty System Analysis
The uncertainty ua2 of the nonlinearity of the system response is recorded as the uncertainty of the nonlinear contribution a2.
The combined uncertainty of the DC quantities VECT,VICT,VST is uv.
The combined uncertainty of the radiation spectra CECT,CICT,CST is uC.
There are uncertainties uεICT, uεST , uTICT, uTST of the high-temperature & the low-temperature calibration radiation source.
Temperature changes of the temperature field inside the instrument will also introduce uncertainty uTR.
Progress & Achievements
Type | Items | Uncertainty Sources
Radiometric Calibration Uncertainty | u(TR,ICT) | System temperature stability and non-uniformity
 | u(TR,ST) | 
 | u(a2) | FTS uncertainty relation
 | u(CICT) | Interferometric efficiency uncertainty, detector temperature stability, measured complex spectral noise
 | u(CECT) | 
 | u(CST) | 
 | u(VICT) | 
 | u(VECT) | 
 | u(VST) | 
 | SCR*RECT | Stray light
Uncertainty of infrared radiation reference source | Temperature u1 | Thermometer calibration, phase change reproducibility, temperature control stability, temperature uniformity, measurement noise
 | Emissivity u2 | Blackbody temperature, heated halo temperature & emissivity, ambient temperature, radiation measurement noise, radiation angle coefficient
 | Ambient temperature u3 | measurement noise
Notes: 本团队针对空间红外辐射基准载荷的辐射测量不确定度进行了详细设计评估，从辐射定标方程的角度进行工程化适配，分解并识别各子项不确定度来源，如非线性、探测直流量、辐射光谱、载荷内部高温定标源和低温定标源的发射率和温度不确定度，仪器内部温度场等

## Slide 14: Phase change cells
Quantitative traceability technology
Miniature phase change gallium fixed point variable temperature space reference black body
14
Design of Miniature Phase Change Cells
Perfusion and encapsulation
Tightness test
Temperature reproducibility measurement
Standard platinum resistance thermometer
Value comparison between Micro Phase Change Cells and National temperature reference
Leakage
Miniature Phase Change Cells Perfusion system
Electron beam welding
Data acquisition system
Temperature controller
The Power Supply
Miniature Phase Change Cells

## Slide 15
Miniature Gallium Fixed Point
There are 4 crucibles filled with gallium and fixed on the bottom of the blackbody.
Designed heating circuit is used to provide stable heat flow into the blackbody.
The temperature of gallium is very stable during the phase change process.
A system of miniature gallium fixed point onboard blackbody radiation source was added on HIRAS of FY-3E satellite.
15
Phase change cells

## Slide 16: Phase change cells
Phase change cells' stability testing has been carried out
Miniature Ga phase change cell
Miniature Ga-In phase change cell
16

## Slide 17: Phase change cells
4 temperature miniature fixed point from 270K to 350K；
With 4 high stable PT；
Measuring the whole range with calibrated PT
Phase change fixed points,
stability ≤10 mK（k=2）
类别 性能 | H2O | Ga-Sn | Ga | 水杨酸苯酯
稳定性（mK） | 3.9 | 1.5 | 2.0 | 1.0
温坪值（K） | 273.339 | 293.779 | 303.262 | 315.137
Temperature Uncertainty of the blackbody ≤84 mK（k=2）
17

## Slide 18: Phase change cells
计算机
Facility
18

## Slide 19: Phase change cells
Material | Al
Diameter of cavity | 44 mm
Diameter of Aperture | 24 mm
Depth | 79 mm
Angle | 60°
E of coating | 0.92
Simulation Emissivity
STEEP3 0.9970
19

## Slide 20: Phase change cells
DATE | BB T（K） | HH T（K） | Diff HH BB（K） | A T（K） | RT T（K） | E
2014/11/6 | 305.09 | 364.34 | 59.25 | 299.25 | 305.40 | 0.9963
 | 303.60 | 379.28 | 75.69 | 299.65 | 304.04 | 0.9955
2014/11/5 | 303.01 | 347.24 | 44.24 | 299.55 | 303.45 | 0.9958
 | 303.39 | 363.31 | 59.92 | 300.15 | 303.87 | 0.9959
 | 305.92 | 381.65 | 75.73 | 300.45 | 306.33 | 0.9958
2014/11/4 | 302.95 | 347.01 | 44.06 | 300.05 | 303.38 | 0.9962
 | 303.22 | 379.06 | 75.85 | 299.85 | 303.70 | 0.9951
2014/10/17 | 301.36 | 361.38 | 60.02 | 295.85 | 301.55 | 0.9967
2014/10/15 | 301.86 | 359.91 | 58.05 | 295.75 | 301.97 | 0.9974
2014/10/14 | 294.33 | 356.93 | 62.60 | 294.75 | 294.46 | 0.9965
平均值 |  |  |  |  |  | 0.9961
标准差 |  |  |  |  |  | 0.07%
10 times measurements during 20 days. The average of E is 0.9961，SD 0.07 %。
Result
20
Notes: 基于以上的实验装置，在20天内测量了10次黑体空腔发射率，发射率的平均值为0.9961，标准偏差为0.07 %。

## Slide 21: Phase change cells
Facility
21

## Slide 22: Phase change cells
Miniature Fixed Point
Fixed points are used to define the ITS-90.
The miniature fixed points as the standard are used to calibrate the thermometers working on the on-orbit blackbody.
22
Platform 2 h
Melting point at 303.113 K with SD 1 mK.
Ga miniature fixed point
The bath keep at 302.75 K
Heat the cavity with 0.1W.

## Slide 23: Phase change cells
The curve of melting at different power
The bath keeps at 302.75 K，the power of cavity heater is 0.066 W，0.0466 W and 0.033W
Ga melting point temperature （111.573 Ω）。
23
Notes: 固定点腔内金属的含量是一定的，相同质量的金属完全相变吸收的能量是相同的，所以加热功率越大金属相变的时间就会越短。同时随着加热功率的增加，黑体的升温速率也会变快，黑体的热惯性变大。
应该使用小加热功率加热黑体保证黑体的升温速率足够小。

## Slide 24
Data colletced in experiment
The momentary temperature when the phase change is finished is called the inflection point.
The inflection point has a relationship with the heating power. For unchanged heating power, the inflection point will have a good reappearence.
In the experiment with the blackbody only, the reproducibility of the inflection point is better than 0.03K.
In the on-orbit mearsurement, inflextion point can be observed clearly. When the heating power is 4.11W, the inflextion point is 304.1K.
The relationship between the heating power and the inflection point is being analyzed with modeling.

## Slide 25
Stability of the thermometer
Annealing furnace
Thermometry bridge
Low-temperature alcohol groove
National benchmarks Water Triple Point
Test Methods: Precision platinum resistance thermometer placed in a temperature annealing furnace (150At ℃) annealing from about2h /Times, by accurately measuring the temperature of the bridge at the triple point of waterResistance thermometerRtp; Then place it into-80Alcohol ℃ of the bath to about annealed 2h /Times, and then measured in the temperature measuring bridge triple point of waterResistance thermometerRtp. Upper and lower limits for the heat treatment after a heating cycle, thermal cycling of the thermometer by multiple triple point of waterRtpPrecision measurements to calculate the stability of platinum resistance thermometer.
data processing: Multiple triple point of waterResistance thermometerRtpMeasurements to calculate the standard deviation.
time | Rtp/Ω | 
 | 2017078 | 2017073
2017.04.18 | 99.95399 | 99.79355
2017.10.19 | 99.94917 | 99.79591
2018.01.11 | 99.95041 | 99.79589
2018.01.19 | 99.94864 | 99.79627
average value/Ω | 99.95030 | 99.79541
standard deviation/mK | 4.8 | 3.1
25
Space Radiometric Benchmark Reference of Infrared Spectral Segment

## Slide 26
System Scheme
Broadband, large field of view infrared interferometer
Broadband, large size & small array infrared sensors
Multi-temperature zone infrared cryogenic optics technology
Efficient refrigeration for 50K temperature technology
Infrared interferometer
Response spectrum: 600 cm-1~2700cm-1
Spectral resolution: 0.5cm-1
Infrared sensors
Circular center symmetry sensor
Effective detection of focal plane without metal reflective layer
Detection rate: 5×1010cmHz1/2W-1@77K
Cutoff wavelength: 17μm@77K, 18μm@60K
HgCdTe Infrared Sensor
26
Space Radiometric Benchmark Reference of Infrared Spectral Segment

## Slide 27
Chip performance has been greatly improved
Detection rate: 5×1010cmHz1/2W-1@77K≈ 1.7 × [3×1010cmHz1/2W-1D@77K(Last year)]
Cutoff wavelength: 17μm@77K
According to the study of the law of HgCdTe performance changing with temperature: 1×1011cmHz1/2W-1 @60K
Cutoff wavelength: 18μm@60K in real test
Progress & Achievements: Broadband, large size & small array infrared sensors
Response spectrum test curve at 77K temperature
27
Infrared sensors

## Slide 28
Progress & Achievements: Interferometer system
Full band
Long wave end
Short wave end
Observe 80℃ and 130℃ black bodies respectively
Response spectrum:
600 cm-1~2700cm-1
Spectral resolution:
0.5cm-1
Performance experiment of interferometer system at room temperature
Frequency stabilized narrow linewidth CO2 laser
Integrating sphere uniform light
Collimator
Infrared sensor
Fourier inversion
28
Space Radiometric Benchmark Reference of Infrared Spectral Segment

## Slide 29
空间红外辐射基准载荷全链路不确定度评估
Uncertainty of Prototype
< 0.085K（k=1）
Space Radiometric Benchmark Reference of Infrared Spectral Segment
29
Notes: 根据红外辐射基准载荷辐射测量合成标准不确定度公式，可以得到载荷的各谱段不确定度结果，如图所示

## Slide 30
谢谢！
www.sitp.ac.cn