## Slide 1: Absolute Intercalibration of Spaceborne Microwave Radiometers
Katherine Wentz, Thomas Meissner, Frank Wentz, and Andrew Manaster
Remote Sensing Systems

## Slide 2
Wentz, K., Meissner, T., Wentz, F., & Manaster, A. (2024). Absolute Intercalibration of Spaceborne Microwave Radiometers. Journal of Atmospheric and Oceanic Technology, 41(12), 1121-1138.
doi: 10.1175/JTECH-D-24-0024.1

## Slide 3: Motivation
The creation of a multidecadal CDR of MW geophysical retrievals requires accurate absolute intercalibration of TBs from multiple operating sensors
GMI is expected to be in orbit at least till 2030. Possibly later if orbital boost is performed.
Adequate follow-on: WSF-M MWI (polar orbit 18:00 LANT).

## Slide 4: Absolute Radiometer Calibration
Determination of calibration parameters: ΔTNL (aNL), η, Cpq
Ideally this would be done pre-launch
In practice this is difficult to do to high accuracy, so post-launch calibration is needed
  Applies particularly to spillover, η. Antenna backlobes are difficult to measure.
Other calibration issues to be dealt with:
  Scan position dependent biases due to intrusion of spacecraft or sensor into Earth FOV
  Hot load thermal gradients due to solar intrusion
  Emissive reflector antenna
  The above anomalies are observed with many sensors (TMI, SSMIS, AMSR-E, WindSat) and are important to correct, but not the subject of this presentation

## Slide 5: TA to TB Equations
Initial TA
Non-Linearity ΔTNL
Non-linearity Correction
Counts
TA
CC
CH
TH
TC
CE
TA,E
T’A = Antenna Temperature with Non-Linearity
CE = Earth Counts
CH = Hot Load Counts
CC = Cold Space Counts
TH = Hot Load Temperature
TC = Cold Space Temperature
aNL = Non-Linearity coefficient, note the ΔTNL term is small and quadratic for WindSat
The Non-linearity correction accounts for the relationship between counts and antenna temperature deviating from linear

## Slide 6: TA to TB Equations
Spillover Correction
TA = Antenna Temperature with Cold Space Spillover and Non-Linearity
T’A = Antenna Temperature with Non-Linearity
TC = Cold Space Temperature
η = Cold Space Spillover into Earth Scene
The Spillover correction accounts for the intrusion of cold space Cosmic Microwave Background (CMB) into the Earth field of view through antenna backlobes where the CMB fraction (η) is a few percent

## Slide 7: TA to TB Equations
Cross-polarization Correction
TBV, TBH = Top of the atmosphere brightness temperature after calibration for V-pol and H-pol
TAV, TAH = Antenna Temperature with Cold Space Spillover and Non-Linearity for V-pol and H-pol
Cp,q = Values for cross-polarization mixing of q into p where p is V-,H-pol and q is V-,H-pol; note, CVH=1-CVV and CHV=1-CHH
The Cross-polarization correction accounts for the mixing of V-pol into H-pol (and vice versa) that is received by the antenna

## Slide 8: GMI TBs are an Absolute Calibration Reference
GMI TBs are absolutely calibrated to within 0.2—0.3 K when compared to the RTM, and are therefore used as a reference TB for other MW sensors
  No vicarious Earth targets required. No specific radiative transfer model (RTM) required.
  4-point Calibration consists of hot load and cold sky view + internal noise diodes
    TC TC + TND TH TH + TND
    Noise diodes check for nonlinearities, fill in for the cold-sky view during orbit maneuvers, monitor the temperature stability of the hot load, and check for intrusions into the cold-sky view
    Pre-launch determination of antenna patterns and non-linearity were accurate
  On-Orbit Maneuvers include (1) deep space maneuver for along-scan bias removal, (2) inverted backlobe maneuver for cold space spillover removal, and (3) nadir look for determining bias offsets in V-pol and H-pol (ability to disentangle biases from cross-polarization effects)
  No calibration anomalies or drifts in GMI TBs (hot load solar intrusion, emissive reflector, …)
  Wentz, F. and D. Draper 2016 paper

## Slide 9: Intercalibration of MW TBs to GMI

## Slide 10: Intercalibration of MW TBs to GMI
The core of the absolute intercalibration methodology is matchups between GMI and other sensors’ TBs [WindSat and AMSR2] over the Open Ocean and Amazon Rainforest
1 hour time window, all data resampled onto 0.25-deg grid; GMI has shifting LANT so can find match-ups with any polar orbiting sensor
Determine calibration parameters for WindSat|AMSR2 in multiple linear regressions that minimize the following cost function
Above, i is the V-pol or H-pol polarization, j is the Open Ocean or Amazon Rainforest observation
We solve for the parameters for each channel separately (10 VH, 18 VH, 23V, 36 VH, 89 VH)
Note, a computationally effective method is to solve for Spillover over Amazon scenes, and then, solve for Non-linearity and Cross-polarization over Open Ocean scenes

## Slide 11: Adjustment to GMI TBs to Account for Different Configuration
Sensor intercalibration method does not depend on a specific RTM
  The method uses an RTM only for a small adjustment to GMI TBs (~ 1 K) in order to match the frequencies and EIA of WindSat|AMSR2
  The TB results are very robust towards changing the RTM or the geophysical input to the RTM

## Slide 12: Calibration Parameter Results: WindSat
WindSat TBs based on prelaunch calibration parameters are ~1-2 K higher over the open ocean and ~3 K higher over the Amazon Rainforest as compared to calibrated TBs

## Slide 13
We used passband center frequencies for the 18H and 23V that differ from the nominal values because the nominal values result in fits for the NL terms that switch sign over the TA range. 18H is 100 MHz lower than nominal value of 18.7 GHz and the 23V is 70 MHz lower than nominal value of 23.8 GHz.
Calibration Parameter Results: AMSR2
JAXA TB (not intercalibrated) – RSS TB Open Ocean |  |  | 
6V | 1.7 K | 6H | 1.7 K
7V | 1.9 K | 7H | 2.3 K
10V | 4.1 K | 10H | 2.8 K
18V | 4.6 K | 18H | 2.9 K
23V | 3.4 K | 23H | 4.3 K
36V | 2.3 K | 36H | 3. 6 K
89V | 2.3 K | 89H | 2.7 K
AMSR2 non-linearities are large and not always quadratic

## Slide 14: Open Ocean TB Calibration Results
GMI Reference
AMSR2
WindSat
TB (calibrated) – TB (RTM) results for Open Ocean scenes
2 V – H mitigates sensitivity to error in atmospheric input (clouds, vapor)
The undulating nature of the TB errors when binned by wind speed is likely a result of uncertainty in the RTM surface roughness model which models the ocean surface emissivity as a function of a fifth-degree polynomial in wind speed
The larger biases in the 89 GHz may be a result of misclassifying clouds in the RTM, which uses cloud profile data from ERA5 (scaled to match GMI); note, the 89 GHz cloud absorption is 2.5x 37 GHz absorption
2017-2019 Matchups

## Slide 15: Sources of Uncertainty
 | Channel | GMI TB Calibration and RTM Uncertainty | AMSR2 TB Calibration and RTM Uncertainty | WindSat TB Calibration and RTM Uncertainty | Ancillary Datasets Uncertainty | NEDT
RMSE | 10 | 0.25 | 0.24 | 0.05 | 0.2 | 0.1
 | 18 | 0.34 | 0.20 | 0.16 | 0.2 | 0.1
 | 36 | 0.05 | 0.48 | 0.28 | 0.2 | 0.1
 | 89 | 0.35 | 1.21 | - | 0.4 | 0.1
Open Ocean Scenes
Monte Carlo simulation is used to determine Ancillary Datasets Uncertainty:
± 0.8 m/s wind speed
± 0.5 K SST
± 1.5 mm water vapor
± 0.1 mm cloud
± 15˚ wind direction
The random uncertainty in TB Calibration contains errors associated with neglecting the dependence of Non-linearity on receiver physical temperature
2017-2019 Matchups, Values in Kelvin

## Slide 16: Amazon Rainforest TB Calibration Results
TB (calibrated) results for Amazon Rainforest scenes
The results for the descending morning passes are better than for the ascending afternoon passes; the ERA5 ancillary fields for SST and clouds/rain that are used in the RTM are more accurate in the local morning time when cloud/rain is less likely and SSTs are more stable
Although it is not shown here, there are large TB biases in the 89 GHz channel likely due to misclassification of clouds in the RTM
GMI Reference
AMSR2
WindSat
2017-2019 Matchups

## Slide 17
 | Channel | GMI TB Calibration and RTM Uncertainty | AMSR2 TB Calibration and RTM Uncertainty | WindSat TB Calibration and RTM Uncertainty | Ancillary Datasets Uncertainty (Mostly Surface Temperature) | NEDT
RMSE | 10 | 0.01 | 0.45 | 0.33 | 1.5 | 0.1
 | 18 | 0.04 | 0.42 | 0.32 | 1.5 | 0.1
 | 36/37 | 0.04 | 0.33 | 0.36 | 1.5 | 0.1
 | 89 | 0.01 | 0.37 | - | 1.5 | 0.1
Amazon Rainforest scenes
MODIS LST vs in-situ data has an RMSE of 1.5 K, which dominates the ancillary dataset uncertainty and the overall standard deviation of GMI, AMSR2, and WindSat errors
2017-2019 Matchups, Values in Kelvin, using Ascending data for AMSR2 and WindSat
Sources of Uncertainty

## Slide 18: Absolutely Calibrated TB Data Available on RSS Webpage
AMSR2 TB Data Available at:
https://data.remss.com/TB/intercalibration/amsr2_TB_maps_daily_025deg/
T. Meissner, K. Wentz and F. Wentz, 2023: RSS AMSR2 TB on daily 0.25 grid, Version 8.2, Remote Sensing Systems, Santa Rosa, CA, https://doi.org/10.56236/RSS-bk
WindSat TB Data Available at:
https://data.remss.com/TB/intercalibration/windsat_TB_maps_daily_025deg/
Meissner et al., Remote Sensing Systems WindSat TB, V08.0, doi:10.5067/WSA80-1CRTB
GMI TB Data Available at:
https://data.remss.com/TB/intercalibration/gmi_TB_maps_daily_025deg/
T. Meissner et al., Remote Sensing Systems GMI TB, V08.2
Available during validation time period: 2017-2019

## Slide 19: Summary and Conclusions
GMI is the best absolute calibrated MW imager
  The inclined orbit and shifting LANT allows match-ups with other MW imaging satellites
  GMI is expected to be in orbit at least till 2030
  Adequate follow-on: WSF-M MWI (polar orbit 18:00 LANT)
Intercalibration of other sensors with GMI
  Performed at TB level: TB Accuracy ~ 0.5 Kelvin or better
  Open ocean + Amazon → Valid over the whole dynamical range of Earth TB
Not dependent on a specific RTM
Backbone of constructing climate data record (CDR) of air-sea essential climate variables (AS-ECV) from intercalibrated satellite microwave imagers
  Wind Speed, SST, Precipitable Water Vapor, Liquid Cloud Water, Precipitation Rate, Sea Ice

## Slide 20: Estimate 18H Non-Linearity to De-Bias Ocean
strange behavior
the NL would need to change sign over the ocean dynamic range
Non-Linearity Correction

## Slide 21
These are the AMSR2 NL form the old Version (V7), which was calibration to the ocean RTM.
The 18H NL did exactly that sign change!
Red Curves are JAXA values for non-linear correction.
Black Curves are values coming from RSS analysis on 2014.

## Slide 22: Ground Truth Validation Wind Speed Retrievals versus Buoys
RMSE: 0.92 m/s
RMSE: 0.87 m/s
RMSE: 0.84 m/s

## Slide 23: Ground Truth Validation Precipitable Water Retrievals versus GPS Stations
RMSE: 1.6 mm
RMSE: 1.6 mm
RMSE: 1.6 mm