## Slide 1: The Status of Microwave Temperature Sounder –III onboard Fengyun-3H Satellite
Juyang Hu,
Xiuqing Hu, Na Xu, Jian Shang, Shengli Wu, Shi Liu
NSMC, CMA

## Slide 2: Outline
Instrument Review
Calibration for FY-3H MWTS-III
Calibration Accuracy Evaluation
Conclusion and Future plan
Cold Wave Processes in the Northern Hemisphere (FY-3H MWTS-III, 2025.10)

## Slide 3: Outline
Instrument Review
Calibration for FY-3H MWTS-III
Calibration Accuracy Evaluation
Conclusion and Future plan

## Slide 4
MWTS: Microwave temperature sounders onboard FengYun -3
FY-3A
FY-3B
FY-3C
FY-3D
FY-3F
FY-3E
FY-3H
2008-2018
2010-2020
2013-2024
2017.11-
2023.08-
2021.07-
2025.09-
MWTS-I
MWTS-II
MWTS-III
MWTS was first carried on the FY-3A satellite launched in 2008.
The newly launched FY-3H is equipped with the third-generation MWTS (MWTS-III).
Seven instruments have been launched, with four currently operating in orbit.
in orbit
1. Instrument Review

## Slide 5
1. Instrument Review
5
CH | Center frequency （GHz） | Band width （MHz） | NEDT （K） | FY-3D NEDT | ATMS | AMSU-A | Metop-SG-A1 MWS
1 | 23.8 | 270 | 0.3 |  | 1 | 1 | 1
2 | 31.4 | 180 | 0.35 |  | 2 | 2 | 2
3 | 50.3 | 180 | 0.35 | 1.2 | 3 | 3 | 3
4 | 51.76 | 400 | 0.3 | 0.75 | 4 |  | 
5 | 52.8 | 400 | 0.3 | 0.75 | 5 | 4 | 4
6 | 53.246±0.08 | 140 | 0.35 |  |  |  | 5
7 | 53.596±0.115 | 170 | 0.3 | 0.75 | 6 | 5 | 6
8 | 53.948±0.081 | 142 | 0.35 |  |  |  | 7
9 | 54.40 | 400 | 0.3 | 0.75 | 7 | 6 | 8
10 | 54.94 | 400 | 0.3 | 0. 75 | 8 | 7 | 9
11 | 55.50 | 330 | 0.3 | 0. 75 | 9 | 8 | 10
12 | 57.290344(fo) | 330 | 0.6 | 0. 75 | 10 | 9 | 11
13 | fo±0.217 | 78 | 0.7 | 1.2 | 11 | 10 | 12
14 | fo±0.3222±0.048 | 36 | 0.8 | 1.2 | 12 | 11 | 13
15 | fo±0.3222±0.022 | 16 | 1.0 | 1.7 | 13 | 12 | 14
16 | fo±0.3222±0.010 | 8 | 1.2 | 2.4 | 14 | 13 | 15
17 | fo±0.3222±0.0045 | 3 | 2.1 | 3.6 | 15 | 14 | 16
Item | MWTS-I | MWTS-II | MWTS-III
Viewing Angle | ±48.3° | ±49.5o | ±53.35o
Samples per scan | 15 | 90 | 98
Step Angle | 6.9° | 2.2° | 1.1°
Nadir Resolution | 50~75 km | 33 km | 33 km
Scan Period | 16s | 5245 ms | 5236 ms
Integration Time | 0.78s | 16 ms | 16 ms
Channels | 4 | 13 | 17
Comparison of the core specifications across the three generations of MWTS
a significant increase in the number of channels, along with comprehensive optimization of key performance metrics such as the scan period and integration time.
These improvements have led to substantial gains in both detection efficiency and accuracy.
Green: MWTS-I channels. Red: The channel addition from MWTS-II to MWTS-III.

## Slide 6
The antenna shield color changes from black to white to improve instrument temperature stability.
SRF standardization brings the measured SRF closer to the ideal box.
The channel crosstalk, in terms of land-sea bias in the sounding channels, is much better than that of the two predecessors. The channel correlation has also been significantly improved.
1. Instrument Review
FY-3H MWTS-III: Characteristics and improvements over predecessors
FY-3H
FY-3E/F
0.1K
0.2K
Sea - land contrast bias
channel correlation (FY-3E/FY-3F/FY-3H)

## Slide 7
All channels meet the requirement.
NEDTs of FY-3H MWTS-III are better than those of FY-3E, except for channel 3.
The NEDTs of channels 1–11 are comparable to that of FY-3F and FY-3H, whereas FY-3H shows slightly lower sensitivity for channels 12–17.
Instrument Noise Status
1. Instrument Review
requirement
Allan Variance

## Slide 8: Outline
Instrument Review
Calibration for FY-3H MWTS-III
Calibration Accuracy Evaluation
Conclusion and Future plan

## Slide 9: Calibration for FY-3H MWTS-III
=
（1）Cold space Radiation Correction
（2）Antenna Pattern Correction
Receiver Calibration：
Antenna Calibration：
（3）Channel Crosstalk Correction

## Slide 10: Calibration for FY-3H MWTS-III
Validation shows that all channels have negative biases.
Cold space observation is contaminated by platform radiation.
The cold space radiation was adjusted based on evaluation result.
The mean bias was improved after the adjustment.
Chart: Channels
| ch5 | ch6 | ch7 | ch8 | ch9 | ch10 | ch11
订正前 | -0.70158 | -0.28473 | -0.30052 | -0.43559 | -0.71895 | -1.5588 | -2.4084
订正后 | -0.46673 | -0.079491 | -0.068535 | -0.15229 | -0.28181 | -0.54642 | -0.90119
Cold space Radiation Correction

## Slide 11: Calibration for FY-3H MWTS-III
Iterate the coefficient for each scan position to minimize the difference between the O‑B at and that of the nadir.
Before and after APC adjustment
Antenna Pattern Correction

## Slide 12: Calibration for FY-3H MWTS-III
Since FY-3C, the MWTS has exhibited channel crosstalk induced by P-1 compression, causing a land-sea bias in the temperature sounding channels, namely the bias over land is lower than that over ocean.
FY-3H MWTS-III has a sea-land bias of 0.12–0.15 K for channels 9–17.
The correction reduces the land-sea bias for all channels to within 0.05 K, while introducing negligible changes to the mean brightness temperature of the channels.
We plan to improve channels 9–11 of the FY-3E MWTS this year.
Channel crosstalk Correction
P-1 compression

## Slide 13: Outline
Instrument Review
Calibration for FY-3H MWTS-III
Calibration Accuracy Evaluation
Conclusion and Future plan

## Slide 14: Calibration Accuracy Evaluation
MWTS-III |  |  | ATMS |  | AMSU-A | 
Center frequency (GHz) | No. | Polarization | No. | Polarization | No. | Polarization
23.8 | 1 | QH | 1 | QV | 1 | QV
31.4 | 2 | QH | 2 | QV | 2 | QV
50.3 | 3 | QV | 3 | QH | 3 | QH
51.76 | 4 | QV | 4 | QH |  | QH
52.8 | 5 | QV | 5 | QH | 4 | QH
53.246±0.08 | 6 | QV |  |  |  | 
53.596±0.115 | 7 | QV | 6 | QH | 5 | QH
53.948±0.081 | 8 | QV |  |  |  | 
54.40 | 9 | QV | 7 | QH | 6 | QH
54.94 | 10 | QV | 8 | QH | 7 | QH
55.50 | 11 | QV | 9 | QH | 8 | QH
57.290344(fo) | 12 | QV | 10 | QH | 9 | QH
fo±0.217 | 13 | QV | 11 | QH | 10 | QH
fo±0.3222±0.048 | 14 | QV | 12 | QH | 11 | QH
fo±0.3222±0.022 | 15 | QV | 13 | QH | 12 | QH
fo±0.3222±0.010 | 16 | QV | 14 | QH | 13 | QH
fo±0.3222±0.0045 | 17 | QV | 15 | QH | 14 | QH
MWTS vs Reference data:
Observe time difference<20min
Distance<15km
Difference of satellite zenith angles < 1°
Two nadir pixels for Ch 1-5
Ten pixels near nadir for others
SNO method: compared with NOAA-21ATMS

## Slide 15: Calibration Accuracy Evaluation
Simulation Study:
(RTTOV+ERA5+MWTS geo parameters)
BT_MWTS (MWTS coefficient) vs BT_MWTS (ATMS coefficient)
analysis BT difference under different zone and different land cover.
BT Comparison Scheme:
window channels 1-6: land pixels, BT compared directly
sounding channels 7-17: sea pixels, BT linear adjustment.
SNO method: SRF difference handling

## Slide 16: Calibration Accuracy Evaluation
For the FY-3H MWTS-III, the mean bias across all channels is below 1 K. The minimum bias occurs at channel 11 (less than 0.1 K), and the maximum bias occurs at channel 3 (approximately 1 K).
The error (both bias and STD) level is similar to that of FY-3F, while being better than that of FY-3E.
The SRF standardization of FY-3H MWTS-III leads to better SNO bias and STD for channels 9–11 than those of FY-3F.
SNO method: compared with NOAA-21ATMS (2025.12.17-2026.01.06)
Mean Bias
Bias STD

## Slide 17: Calibration Accuracy Evaluation
Simulation : using RTTOV+ERA5
Samples: 60°S ~60°N, sea pixels for all scanlines
After APC coefficient adjustment, the O-B differences across all 98 scan positions for all channels are less than 0.2 K.
O-B method: Along Scan bias

## Slide 18: Calibration Accuracy Evaluation
Samples: 30°S ~30°N, sea pixels, ten pixels nearby nadir
Channels 7–13 are approximately 0.1 K, channels 6 and 14–17 range from 0.15 K to 0.2 K; window channels 1–5 are between 0.5 K and 1 K.
Different satellites exhibit seasonal variations in their A/D biases. FY-3H has shown a higher bias in descending orbits than in ascending orbits over the past six months.
O-B method: Ascend/Descend bias
Ch9
Ch10
Ch11
Ch3
ATMS
FY-3E
FY-3F
FY-3H

## Slide 19: Outline
Instrument Review
Calibration for FY-3H MWTS-III
Calibration Accuracy Evaluation
Conclusion and Future plan

## Slide 20: Conclusion and Future plan
FY-3H MWTS-III has several improvements, i.e. SRF standardization, channel crosstalk.
The platform environment of FY-3H is worse than that of FY-3F, with obvious radiation interference during cold space viewing and Earth-viewing observations.
The calibration accuracy of FY-3H is significantly better than that of FY-3E, and similar to that of FY-3F.
Prepare for FY-3 (10) satellite (will be launched this year).
Calibration quality improvement for FY-3 MWTSs, i.e. striping noise removing, orbital bias correction.

## Slide 21: Thank you for your attention!
Email：hujy@cma.gov.cn