## Slide 1: 2026.03.24
Two-Point Calibration and Nonlinearity Adjustment Optimization of FY-3D MWRI
Shi Liu, Shengli Wu
National Satellite Meteorological Center
(National Center for Space Weather)

## Slide 2: FY-3D exhibits drift in TBs due to blackbody temperature monitoring issues.
GMI data are relatively stable, and nonlinear coefficient correction using one-year-long data can be applied to inter-calibrate FY-3D to the GMI standard and applied to operational applications.

## Slide 3: A systematic four-phase framework for dynamic bias elimination
Data Preprocessing and Matching up
Phase1
Phase4
Phase2
Phase3
Action:
Spatial/temporal geometry matching
Metric:
Space < 0.03°
Time < 10 mins.
Radiation Reference Alignment
Dynamic Error Compensation
Bias
Elimination
Action: Double Difference via
ERA5/MWSTRM4.
Metric:
Sea/Rainforest target isolation.
Action:
Isolating nonlinearity & calculating nonlinear coefficient.
Metric: Extracting µcold intensity.
Action:
Residual PDF shifts & dataset generation.
Metric:
Zero-centered residuals.

## Slide 4: Rigorous SNO matching guarantees statistically uniform global datasets
SNO selection criteria: time difference less than 10 minutes, spatial difference less than 0.03 degrees, and the difference between the 3×3 and 7×7 luminance temperature STD of each channel less than 1.
The data throughout the year is evenly distributed globally, which meets the calibration requirements.
Homogeneity Screening
Global Distribution

## Slide 5: Radiation Reference Alignment via Double Difference (DD)
Observed differences
Background simulations
(ERA5 + MWSTRM4 RT model).
The hot-brightness temperature is derived from the DD of Tropical rainforests without polarization.
The sea-brightness temperature is derived from the DD of clear-sky and moisture-free ocean conditions.
Notes: 热点附近的窗口散点只用于“检查一致性/合理性以及温度分布”，不是用来拟合热点值。

## Slide 6
 | Sea-Scene |  |  | Rainforest-Scene |  | 
Channel | Mean Calibration differences (K) | STD of Calibration differences (K) | Mean GMI TB (K) | Mean Calibration differences(K) | STD of Calibration differences(K) | Mean GMI TB (K)
10V | 5.9355 | 1.0916 | 167.4522 | 2.3293 | 1.0918 | 285.5751
10H | 7.7473 | 1.2228 | 90.1329 | 1.9578 | 1.3980 | 284.1416
18V | 1.7896 | 1.0313 | 190.8271 | 0.2240 | 0.9896 | 286.5534
18H | 1.9594 | 1.0252 | 122.8313 | -0.7433 | 1.1005 | 285.5253
23V | 1.5038 | 1.0070 | 218.0963 | 1.6656 | 0.8534 | 287.6833
36V | 6.0545 | 0.9052 | 213.6787 | 3.4869 | 0.6691 | 285.5545
36H | 6.2381 | 0.8484 | 151.1619 | 3.5990 | 0.9224 | 284.7570
89V | 1.8359 | 0.9000 | 258.3129 | 1.7922 | 0.7618 | 288.3535
89H | 1.5266 | 1.0643 | 222.8525 | 1.5648 | 0.8362 | 287.9724
The determined sea and Rainforest scene deviation values by the Double Difference

## Slide 7: Obtain Deep Space Observations and Determine the two-point linear calibrations
Deep space observations can be calculated using the following formula:
which
​ corresponds to the physical temperature of the cold Reflector,
thus determining the cold point.
All cold scene and hot scene sample points represent non-linear characteristics on both sides of the straight line between the cold and hot points.
𝑎𝐻 𝑐𝑜𝑟𝑟𝑒𝑠𝑝𝑜𝑛𝑑𝑠 𝑡𝑜 𝑡ℎ𝑒 𝑒𝑚𝑖𝑠𝑠𝑖𝑣𝑖𝑡𝑦 𝑜𝑓 𝑡ℎ𝑒 𝑐𝑜𝑙𝑑
Reflector, 𝑤ℎ𝑖𝑐ℎ 𝑖𝑠 𝑎𝑚𝑜𝑛𝑔 𝑡ℎ𝑒 𝑡𝑒𝑛 𝑐ℎ𝑎𝑛𝑛𝑒𝑙𝑠：
0.9996,0.9996,0.9995,0.9995,0.9995,
0.9995,0.9993,0.9993,0.9990,0.9990

## Slide 8: Isolating instrument nonlinearity across the ocean cold scene
as a function of observed brightness temperature
Counts from the MWRI
calibration linkage.
Localized constant representing channel nonlinear intensity
In particular: the gain coefficient , the cold-end calibration count , the hot-end calibration count , and the observation count all come from the MWRI calibration/voltage chain.

## Slide 9
Distribution of residual probability density functions after applying linear correction and subsequent correction to the nonlinear coefficients on all channel samples.
Systematic bias elimination yields normalized,
zero-centered residuals
STATE 1: RAW. Biased, scattered, and structurally skewed prior to calibration.
STATE 2: LINEAR CALIBRATION. Shifted toward zero via Double Difference, but retaining nonlinear coefficients skew.
STATE 3: NONLINEAR CORRECTION. Perfectly zero-centered, normalized Gaussian distribution. Nonlinear artifacts eliminated.

## Slide 10
Residuals as a function of observed brightness temperature
Inter-satellite residuals demonstrate global consistency across all channels
Observation Range：The red line (Nonlinear Corrected) remains tightly bounded near 0 K across the entire 150K to 300K spectrum.
Significance：The dynamic error compensation successfully flattens the response curves, proving that the FY3D-MWRl observations now reliably mirror the GMI reference at all global observation temperatures.

## Slide 11
FY3D-MWRI dataset is fully calibrated, stable, and ready for operational application
High-quality SNO datasets matched-up and spatially homogenized
Radiation Reference Alignment unified via DD.
Nonlinearity dynamically compensated
Inter-satellite systematic biases calibration
Core Conclusion:
The updated cross-calibration parameters have successfully eliminated systematic nonlinear biases. Currently, FY3D-MWRI remote sensing data has achieved seamless integration with the GMI global reference standard.

## Slide 12
THANKS