## Slide 1: 8 years TROPOMI long-term instrument degradation derivation, validation and lessons learned
Erwin Loots 1 , Edward van Amelrooy 1, Mirna van Hoek 1, Emiel van der Plas 1, Nico Rozemeijer 2
1KNMI, Royal Netherlands Meteorological Institute, De Bilt, The Netherlands,
2TriOpSys B.V., Utrecht, The Netherlands

## Slide 2: TROPOMI on-board Sentinel-5 Precursor
Single payload
Hyperspectral imager with 4 spectrometers
Sun synchronous orbit (MLTAN 13.30)
Pushbroom with ~ 2600 km swath
High spatial sampling (down to 5.5 km x 3.5 km)
Daily measurements of the Sun via Sun port
Launched 10/2017
Nominal operations since 30/4/2018
TROPOMI spectral bands – based on calibration data |  |  |  |  |  |  |  | 
Spectrometer | UV |  | UVIS |  | NIR |  | SWIR | 
Band ID | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Spectral range [nm] | 267-300 | 300-332 | 305-400 | 400-499 | 661-725 | 725-786 | 2300-2343 | 2343-2389
Spectral resolution [nm] | 0.45 - 0.5 |  | 0.45 - 0.65 |  | 0.34 - 0.35 |  | 0.227 | 0.225
Spectral sampling [nm] | 0.065 |  | 0.195 |  | 0.126 |  | 0.094 | 
Spatial sampling [km2] | 5.5 x 28 | 5.5 x 3.5 | 5.5 x 3.5 |  | 5.5 x 3.5 |  | 5.5 x 7 | 
Detector binning factor | 16 | 2 | 2 | 2 | 2 | 2 | 1 | 1

## Slide 3
Tropomi long-term stability
Tropomi long-term stability
Tropomi long-term stability
We want to know the instrument degradation for radiance and irradiance measurements (and hence reflectance degradation)
Goal: acquire Calibration Key Data (CKD) such that long-term instrument drift of both radiance and irradiance are corrected for by the L01b Processor
The only (more or less stable) in-flight measurements we have are:
solar measurements: both daily (QVD1 diffuser) and weekly (QVD2 diffuser)
After the derivation, we can try to validate the L1b products by:
Using ground-based locations suitable for monitoring the long-term stability of the sensor (e.g. PICS)​
Comparative studies using various satellite instruments with similar characteristics covering the same spectral range​
Comparisons with radiative transfer model calculations using reliable input parameters (cloud- and aerosol-free cases only)
What about alternatives… e.g. OMI approach (radiance derived from Antarctic IF-measurements)…...

## Slide 4: Instrument degradation (1/3)
Goal: Determine (and correct for) the instrument degradation
Diffusers (QVDs) degrade; QVD1 (main, daily) six times faster than QVD2 (backup weekly), proportional to exposure.
The instrument excluding the diffusers (mainly spectrometer) also degrades, proportional to in-orbit time.
 Separate diffuser-only degradation and common/instrument degradation
Degradation strongest for short wavelengths
Spatial/spectral smoothness not prescribed/modelled, but is emerging from the data

## Slide 5: Instrument degradation (2/3)
(Simple) example:
Consider daily and weekly irradiance measurements on 01.05.2023, at 360nm
Use that TROPOMI week is six days, thus QVD1 degrades six times faster than QVD2.
S_qvd1: - 9%; S_qvd2: -4%
Solution: common/instrument degradation: Degr_instr = ?? ;
QVD1-only degradation: Degr_qvd1 = ?? ;
QVD2-only degradation: Degr_qvd2 = ??

## Slide 6: Instrument degradation (2/3)
(Simple) example:
Consider daily and weekly irradiance measurements on 01.05.2023, at 360nm
Use that TROPOMI week is six days, thus QVD1 degrades six times faster than QVD2.
S_qvd1: - 9%; S_qvd2: -4%
Solution: common/instrument degradation: Degr_instr = 3% ;
QVD1-only degradation: Degr_qvd1 = 6% ;
QVD2-only degradation: Degr_qvd2 = 1%

## Slide 7: Issue 2: Instrument degradation (2/3)
(Simple) example:
Consider daily and weekly irradiance measurements on 01.05.2023, at 360nm
Use that TROPOMI week is six days, thus QVD1 degrades six times faster than QVD2.
S_qvd1: - 9%; S_qvd2: -4%
Solution: common/instrument degradation: Degr_instr = 3% ;
QVD1-only degradation: Degr_qvd1 = 6% ;
QVD2-only degradation: Degr_qvd2 = 1%
→ RAD: -3%;
→ IRR:-9%;
→ REFL: +6%

## Slide 8: Instrument degradation (3/3)
Formalization: For all irradiance msmts:
Degr S_qvd1 = D_qvd(t_qvd1) * D_instr(t_orb) * H_i
Degr S_qvd2 = D_qvd(t_qvd2) * D_instr(t_orb) * H_i * K_i
(diffuser degr) (instrument degr) (jitter)
We need models for degradation in time,
e.g. D_instr(t) = exp(-at)
Note: engineering data provides cumulative
diffuser usage times t_qvd1, t_qvd2 per orbit
(Degradation in time, as observed from
the measurements, is not smooth -->
→ this is the so-called irradiance jitter)
2018
2018
2025
2025

## Slide 9: Intermezzo: irradiance jitter
From orbit-to-orbit, small variations (smooth in spectral & spatial direction) are observed. No correction possible (unless, maybe, from hindsight)
This ‘irradiance jitter’ is correlated VIS-NIR and UV-SWIR. Source: diffuser mechanism?
→ Any degradation model cannot be more accurate than these undeterministic variations (giving a lower limit to model residuals)
This disappointing irradiance jitter, once accepted, can help in assessing the quality of our degradation modeling
24 March 2026
GSICS Annual Meeting 2026 - TROPOMI L1b UVN
2025
2018
2018
2025

## Slide 10: Solving linear systems (1/3)
Formalization: For all irradiance msmts:
(eq1) Degr S_qvd1 = D_qvd(t_qvd1) * D_instr(t_orb) * H_i
(eq2) Degr S_qvd2 = D_qvd(t_qvd2) * D_instr(t_orb) * H_i * K_i
Per week, [S_qvd1, S_qvd2, t_orb, t_qvd1, t_qvd2] available
First model: D_instr(t) = exp(-at), D_qvd(t) = exp(-bt)
Method:
i) eq1/eq2 gives linear system ;
solving gives D_qvd and residual 1/K_i
ii) substitute D_qvd in eq1, gives linear system;
solving gives D_orbit and residual H_i
Clearly, we have lack-of-fit: the
residuals are much larger than the
irradiance jitter.
A less simple model is needed….
24 March 2026
GSICS Annual Meeting 2026 - TROPOMI L1b UVN

## Slide 11: Solving linear systems (2/3)
Formalization: For all irradiance msmts:
(eq1) Degr S_qvd1 = D_qvd(t_qvd1) * D_instr(t_orb) * H_i
(eq2) Degr S_qvd2 = D_qvd(t_qvd2) * D_instr(t_orb) * H_i * K_i
Per week, [S_qvd1, S_qvd2, t_orb, t_qvd1, t_qvd2] available
More extensive model: D_instr(t) = exp(-at – bt^2 – ct^3 -dt^4 ), D_qvd(t) = exp(kt – lt^2 - mt^3)
Method:
i) eq1/eq2 gives linear system ;
solving gives D_qvd and residual 1/K_i
ii) substitute D_qvd in eq1, gives linear system;
solving gives D_orbit and residual H_i
This model ‘looks’ better: no
lack-of-fit, the residuals are
similar as the irradiance jitter
while the jitter itself is not part of
the solution (no overfitting)
24 March 2026
GSICS Annual Meeting 2026 - TROPOMI L1b UVN

## Slide 12: Solving linear systems (3/3)
Formalization: For all irradiance msmts:
(eq1) Degr S_qvd1 = D_qvd(t_qvd1) * D_instr(t_orb) * H_i
(eq2) Degr S_qvd2 = D_qvd(t_qvd2) * D_instr(t_orb) * H_i * K_i
Per week, [S_qvd1, S_qvd2, t_orb, t_qvd1, t_qvd2] available
Thus, we obtain different solutions for D_instr CKD (red line)
and D_qvd CKD (black line), depending on our modeling.
Worse: before even attaining overfitting (by imposing even higher-order polynomials) we encounter numerical stability problems: bad ‘matrix condition numbers’...
Worse: (Matrix condition can partly be improved by tools like orthogonal polynomials.) The t_qvd1 and t_qvd2 times do not overlap much, due to very extensive E1 in-flight calibration of the QVD1 diffuser. This means that the linear system for D_qvd is ill-posed.
However, monitoring the residuals vs jitter helps in assessing the quality of the solution, and hence of the degradation CKDs.
24 March 2026
GSICS Annual Meeting 2026 - TROPOMI L1b UVN
Unstable solution
Stable solution
Lack-of-fit
Structure in residuals...
Seems OK….
Ockham’s razor?
Structure?

## Slide 13: Degradation CKDs: on uniqueness
The details of the modeling of diffuser degradation and instrument (common) degradation directly determine the CKDs and therefore the relation between IRR and RAD degradation:
RAD measurements: are corrected with instrument-degradation CKD
IRR measurements: are corrected with instrument-degradation CKD AND QVD1-degradation CKD.
We know the sum of the two (the irradiance msmts themselves), we try to find the the parts.
Assumption: Degr_qvd is smoother than Degr_instr.
No uniqueness guaranteed. Modelling involves fitting smooth curves through daily and weekly msmts
Lack-of-fit/underfitting: residuals too large.
Overfitting: residuals (jitter) become part of CKDs.
Numerical instabilities: The linear systems to be solved have large condition numbers. Check the residuals (jitter): they should be spectrally and spatially smooth.
Irradiance jitter: from disappointment to acceptance to powerful indicator

## Slide 14: Why it matters
Question: Why this cumbersome approach of disentangling the instrument-without-diffuser degradation component from solar (irradiance) measurements?
Why not looking directly at at Earth (radiance) measurements?
Answer: Direct derivation could be tried
-using massive sampling of global radiance measurements. Problem: invariants?
-using ground-based locations (Sahara pseudo-invariant calibration sites PICS) or CEOS LANDNET sites, e.g. Railroad Valley
… but let’s first look at validation with these methods:
24 March 2026
GSICS Annual Meeting 2026 - TROPOMI L1b UVN

## Slide 15: Validation (1/3): massive sampling of Earth radiance
24 March 2026
GSICS Annual Meeting 2026 - TROPOMI L1b UVN
Idea: to derive trends from massive amounts of measurements, either directly (radiance in southern hemisphere/Pacific) or indirectly (AAI from specific wavelength ratios)
Problems: trends through seasonal variations, large sensitivities, AND: on Earth, which quantity (if any) is invariant?

## Slide 16: Validation (2/3): PICS
24 March 2026
GSICS Annual Meeting 2026 - TROPOMI L1b UVN
From: https://mpc-vdaf.tropomi.eu
Stability of reflectance (B3,4,6) has been obtained from near cloud-free, near nadir measurements over 20 Pseudo-Invariant Calibration Sites located in the Saharan Desert and the Arabian Peninsula from May 2018 through January 2026. For all bands, the long-term drift is statistically not significant for the bulk of the sites.

## Slide 17: Validation (3/3): Railroad Valley
24 March 2026
GSICS Annual Meeting 2026 - TROPOMI L1b UVN
400nm
VZA < 10deg
CF < 0.2
440nm
VZA < 10deg
CF<0.2
Tropomi:
-TOA reflectance from SCLNLER product, 1nm resolution ) regridded on 0.125 deg ^2 resolution
-330-790nm, with SZA/VZA and cloud fraction information
-Selections on near nadir and low cloud fractions
RRV: (RRV Playa):
-TOA reflectances from RadCalNet archive
- 400-2500nm (10nm resolution)
Number of useful data points is very low. The statistics is heavily influenced by the choice of allowed ground pixels/data filtering.
Nevertheless, this is considered a successful comparison

## Slide 18
Resume: Long-term degradation of Tropomi
Using solar measurements with two diffusers, CKDs are derived for the long-term degradation of both radiance and irradiance
Based on a few assumptions, an extensive model leads to a linear system, to be solved with some mathematical and instrumental caveats.
By doing this, many lessons learned: on instrument behaviour, and unexpected findings (straylight, residuals, solar line index timeseries...)
Validation is done using ground-based measurements (Sahara PICS, RRV). No clear remaining trends (which is good), but low spatial/spectral resolution and limited accuracy (less good).
Large-scale sampling comparisons remain inconclusive
Therefore: for Tropomi, calibration using the Sun remains the best option
On numbers: Total (irradiance) degradation (Band 1) is < 22% after eight years, and slowing. No problems concerning extension of mission lifetime.

## Slide 19: Thank you!
general information: http://www.tropomi.eu/
more on TROPOMI calibration: Kleipool et al. 2018 https://doi.org/10.5194/amt-11-6439-2018
Ludewig et al. 2020 https://doi.org/10.5194/amt-2019-488
Acknowledgement: Sentinel-5 Precursor is a European Space Agency (ESA) mission on behalf of the European Commission (EC). The TROPOMI payload is a joint development by ESA and the Netherlands Space Office (NSO). The Sentinel-5 Precursor ground-segment development has been funded by ESA and with national contributions from The Netherlands, Germany, and Belgium.
GSICS Annual Meeting 2025 - TROPOMI L1b UVN