## Slide 1
Brad Fisher, Glen Jaross, Sri Madhavan, M. K. Rama Varma Raja, Jeremy Warner, Mike Linda & OMPS L1B Team
Calibration of the NOAA-20 OMPS Instrument for NASA Products

## Slide 2
NOAA –N20 Launched on November 18, 2017
Level 1A
Uncorrected, unprocessed signals
Time-ordered counts (uncalibrated)
Conversion of Housekeeping Data
Level 1B
Radiance (calibrated)
Irradiance (calibrated)
Geolocation
Spectral
LEVEL 1 PRODUCTS
Sensor | Nadir Profiler (NP) | Nadir Mapper (NM)
Spectrometer Type | Double-grating spectrometer | Grating Spectrometer
Spectral range (nm) | 248-312 | 300-423
spectral sampling interval (nm) | 1.1 | 1.1
Band Pass (nm) | 0.42 | 0.42
Pixel Spatial Resolution (km2) | 50 x 50 | 17 x 13
Cross Track Ground Pixels | 1 | 140

## Slide 3
Flowchart illustrating the Earth View (EV) Level 1B generation process
NASA Level 1 Processing
Convert housekeeping details to physical variables (P, V, I, etc.)
Wavelength Assignment
Geolocation
Subtract Bias, Dark, Smear Signals
Convert EV Signals to RAD
Stray Light Correction
Cal. Table Input
Data Flagging (QC)
L1B
L1A
On Board Calibration (non-linearity, pixel gain correction, and
binning)
Corrected Counts
DATASET | DATASET DEFINITION
CBC | Band Center Wavelengths of each CCD pixel
SRG | Elevation angles at the center of each CCD pixel
FOV | Elevation slit function for each CCD pixel
RAD | Radiometric Coefficients for each CCD pixel
STB | Sample Table
Notes: RADIANCE COEFFICIENTS (CALCONST)
IRRADIANCE COEFFICIENTS (SIRR-CC)
SDR-CC (contains the mounting matrix and various flags)

## Slide 4
NOAA-20 Day One Radiometric Calibration
Fitting to a low order polynomial reduces test artifacts and fine spectral structure
January 28, 2018, is defined as Day One, when the first solar measurements using the reference diffuser are made
All subsequent radiance calibration adjustments are applied relative to this date
Day One EV BCW wavelength registration and intra-orbital wavelength variations are characterized
BSDF adjustments are based on Antarctic ice radiances at 354 nm – provides radiometric consistency with other instruments (OMPS and OMI)
Pre-Launch Calibration

## Slide 5
Level 1B pixel quality flagging: pixel affected by transients due to high energy particle hits
Principal Component Analysis is used to identify transients in the EV image data (on corrected counts)
Use set number of PCs to define the background counts
2) Reconstruct the background counts
3) Compute difference between observed and reconstructed counts and apply a count threshold.
Quality Flags (Transients)
Global map of NP transient monthly detections

## Slide 6
Stray Light Correction
  Stray Light (SLT)
    An SL model is derived from a pre-launch characterization of stray light with updates from inflight validations and applied via a 4-D correction Jacobian.
    Stray light is calculated and removed event-by-event
NM Calculated Stray Light % and Residual % for a single event
NP Calculated Stray Light % and Residual % for a single event
NP
NM

## Slide 7
Wavelength Assignment
Intra-Orbital Wavelength Assignment
Wavelength calibration is performed with the solar flux measurements by making use of the spectral structure of the solar Fraunhofer lines compared to a Day One reference spectrum
Wavelength assignment accounts for the seasonal and long-term wavelength shifts, and for NM wavelength assignment also accounts for intra-orbital shifts
NM
~0.01 nm shift over 6 years
OMPS N20 NP Long-Term Spectral Shift
OMPS N20 NM Intra-orbital spectral shift
NP
Long-term Shift ~0.012 nm
Spectral Shift [nm]

## Slide 8
Solar Activity Correction Mg II Index
The Mg II Index is computed every day from EV observations and used to correct for changes in solar activity during the mission relative to Day One
SA Corrections only applied to NP
Day One
NOAA-20 NP EV
Aura OMI
Mg II Index Comparison: NOAA-20 OMPS NP EV and Aura OMI
Mg II Index [core/wing]

## Slide 9
NOAA-20 NP and NM Percent Change per Year in Irradiance (%)
Working diffuser degradation
Change in reference irradiance (instrument change)
NP and NM working diffuser degradation range ~ -0.1 to -0.4% per year
NP instrument changes ranges between -0.3 and -0.1 per year
NM instrument changes is about 0.1% per year

## Slide 10
Derivation of Time-Dependent Calibration (CF) of NP
The calibration factor (CF) accounts for sensitivity changes in each CCD pixel over time
Spectral Pixels
Spatial Pixels
N20 NP CCD image of Operational ACF (CF)
Reference Events 2025-04-11
Increasing Wavelength
The lowest wavelength in the image region is 248 nm and the highest is 312 nm
The NP instrument change is about 2%

## Slide 11
Derivation of Time-Dependent Calibration (CF) of NM
Spectral Pixels
Increasing Wavelength
Spatial Pixels
N20 NM CCD image of Operational ACF (CF)
Reference Events 2025-04-11
The lowest wavelength for NM valid image region is around 300 nm and the highest is around 423 nm.
The NM instrument change is about 1%
Example of a NM CCD pixel at 347 nm shows sensitivity change over the mission
predicted

## Slide 12
BCW wavelengths for L2 retrievals
X = NM
X = NP
N20 Residuals [% change]
RTM Assessment of OMPS NOAA-20 Performance
for Clear and Cloudy Conditions using OMPS-NPP Ozone
Clear Sky Case (REF < 0.15)
DC Case (ECF >=0.8)
PC Case (0.3 < ECF < 0.5)
NPP and NOAA-20 fly in parallel orbits at the same latitude separated by 12º longitude
A Radiative Transfer Model is used to assess performance based on OMPS-NPP Ozone profile and OMPS-N20 viewing conditions (geometry, surface, and cloud conditions)
For Clear and DC cases: The RTM is run and then the NOAA-20 reflectivity is recalculated before computing normalized radiances for Clear and Deep Convections (DC) cases.
For partly cloudy (PC) skies: Apply MLER cloud model and then run RTM twice for clear (REF = 0.15) and cloudy (REF = 0.8). Use the 331 residual to adjust the whole NP and NM wavelength spectrum
N20 Residuals [% change]
Apply MLER Cloud Model
N20 Residuals [% change]
BCW wavelengths for L2 retrievals
X = NM
X = NP
BCW wavelengths for L2 retrievals
X = NM
X = NP

## Slide 13
Monthly Zonal Means of Residuals
The pixel sampling is constrained to a geographical region between 30S and 30N
Clear sky cases are uniformly distributed but mostly over the tropical oceans
Extending the analysis to cloudy cases (DC and PC) would increase the sampling while also providing some sampling over land
This analysis shows consistency between two OMP instruments and can be used to identify calibration issues and for applying soft calibration to the Level 2 ozone products
Clear DC PC
Sampling (clear, DC, PC)
MZM Clear Sky Case
Sampling (clear)
RTM does not account for Raman Scattering

## Slide 14
Summary
Key Improvements to the NOAA-20 Calibration
Improvement to prelaunch tables (BPS, CBC, IRD, RAD, etc.)
Time-dependent wavelength assignment with intra-orbital adjustments
Time-dependent calibration factor to account for sensor throughput change
Extended NM range from 380 to 423 nm
Geolocation error fixed
Improved stray-light and spectral-ghost correction
BSDF adjustments based on Antarctic ice radiances
New Quality flags for edge-pixels, transients, RTS behavior
RTM residual analysis produced good results for clear and cloudy conditions
We plan to use the RTM residual analysis to apply a static calibration adjustment in our L2 ozone product generation.
Expected Official Release version 2 of N20 L1B Data to GES DISC: Summer 2026