## Slide 1
Improving Mission-Long Quality Stability of SNPP OMPS NP SDR Data through Solar Activity and Optical Throughput Degradation Adjustments
Steven Buckner1
NOAA/NESDIS/STAR OMPS SDR Team
NOAA Affiliate
March 24, 2026
GSICS Annual Meeting
UVN Subgroup
Co-Authors: Banghua Yan2, Trevor Beck2, Xin Jin1, and Lawrence Flynn2
1ERT, 2NOAA/NESDIS/STAR

## Slide 2: Outline
Background
  OMPS instrument
  Solar flux measurements and sensor optical throughout degradation
Analysis of SNPP OMPS NP Solar Flux Adjustments and Instrument Throughput Degradation
  Computation Method
  Wavelength shift correction
  Solar activity correction
  Instrument Throughput Degradation Correction
Overview of Preliminary Results Via Reprocessing
Summary and Path Forward

## Slide 3: Background on OMPS
Ozone Mapping and Profiler Suite
Currently onboard both Suomi-National Polar Orbiting Partnership (SNPP), NOAA-20 and NOAA-21 satellites
Includes 3 different sensors
  Nadir Profiler
  Nadir Mapper
  Limb Profiler (not included on NOAA-20)
Coverage of UV wavelengths between 250 and 380 nm
  NP covers 250 to 310 nm
  NM covers 300 to 380 nm
Image from Rault & Loughman 2013, courtesy of BATC

## Slide 4
Each OMPS NP instrument contains two solar diffusers
  Working diffuser: takes solar irradiance measurements once every two weeks
  Reference diffuser: takes solar irradiance measurements once per year
Repeated exposure to solar irradiance causes the diffusers to degrade slowly over time
Optical throughput degradation beyond the diffuser affects both the solar and Earth measurements
The amount of optical throughput degradation experienced by an instrument exclusive of the diffusers can be determined by scaling and combining the measurements from both of the diffusers
  Assumes that the diffuser degradation rates are the same per exposure
Degradation rates are impacted both by wavelength shift and by solar activity
Neither degradation nor solar activity have been corrected for in the past OMPS NP Sensor Data Record (SDR) processing
Monitoring OMPS Sensor Optical Throughout Degradation
Through Two Diffusers’ Solar Flux Measurements
Examples of working diffuser solar flux measurements (relative to day-1) for three channels. Degradation rates and impacts from solar activity and wavelength shift vary based on the channel.

## Slide 5: OMPS Sensor Optical Throughput Degradation Computation Method
Calibration of Earth Radiance at pixel ( is expressed to be
where (1) is the radiance calibration coefficient from pre-launch measurements;
(2)
is the corrected radiometric counts of earth view radiance
Calibration of solar irradiance is expressed to be
F
where (1)is the solar calibration coefficient from pre-launch measurements;
(2)
(3) is the solar diffuser plate reflectivity change
is the corrected radiometric counts of solar irradiance
The degradation rate of the instrument can be described through
  and are arbitrary functions that fit the reference and working diffusers, respectively
  is a correction factor that is intended to account for differences in the optical properties of the diffusers
    Due to a lack of dedicated data, and that the diffusers were designed to have similar characterization, is assumed to be 1.
  is the exposure time for each diffuser
    It is assumed that the rate of degradation is consistent for a given number of measurements and that the time it takes to make each measurement is consistent, so can be converted to , the ratio of number of measurements for each diffuser
    The original equation can be rewritten as:
Equation from OMPS ATBD and C. Pan et. al.
More information can be found in the OMPS ATBD
()

## Slide 6
The solar signals, in addition to degradation related to exposure time, are also dependent on solar activity and wavelength shifts
Before being able to calculate degradation rates, wavelength shift and solar activity must be accounted for through pre-processing
Wavelength shift
  Determined based on the movement of the center of the 280nm absorption feature
  Uses the 5 wavelength channels closest to the 280nm minimum
  A curve fit is used to find the location of the minimum of the curve
  The difference between the measured and calculated minima are determined
  Varies over time, following an annual pattern
OMPS Solar Measurements:
Impacts of Wavelength Shift and Solar Activity
Shift (nm)
Shift (nm)
(top) Examples of lifetime wavelength shift values for SNPP and NOAA-20
(bottom) Example of calculation of wavelength shift for SNPP
Measured Solar Flux
Curve Fit

## Slide 7
Solar UV irradiance can experience large variation due to solar activity, thus leading to change of solar irradiance in magnitude that are measured by satellite instruments such as OMPS NP (Deland and Cebula, 1993)
Solar activity can cause variations of solar irradiance, particularly in the NP wavelength range, also leading to changes in Earth-view (EV) radiance in the 3 NP instruments
  Solar activity is variable on both a 27-day and 11-year cycle
In the UV bands, one way to describe the solar activity is to use Mg-II core-to-wing index (Heath and Schlesigner, 1986)
The Mg-II absorption feature is centered around 280nm. The Mg-II Index can be calculated by taking the ratio between the Mg-II core and the wing continuum around it. This is done by taking an average of the two wing values. A curve fit is used for both the core and wing to have more accurate positions
Correction for Solar Activity (1/3):
Overview
Image from Deland and Cebula, 1993

## Slide 8
Correction for Solar Activity (2/3):
Calculation of Mg-II Index Values
Mg-II Index values were calculated for the lifetime of each of the OMPS NP sensors using RawFlux solar data
  RawFlux solar data was produced by the NASA OMPS team and obtained through the NASA SIPS System
These values represent an approximation of changes in solar activity over time and are used to calculate the solar activity scale factors
To calculate scale factors, Mg-II Index values must be taken from large, short-term changes in Mg-II Index values
  GOME2 daily Mg-II Index values were used as a reference to find dates that lined up approximately with 27-day solar cycle maxima and minima

## Slide 9
Correction for Solar Activity (3/3):
Calculation of Solar Activity Scale Factors
Once several pairs of consecutive measurements with large Mg-II Index changes are found, they can be used to calculate scale factors
Scale factors are calculated using percent change in Mg-II Index and percent change in spectral value for these matched dates
  A line is fit to a scatter plot of these values over time. The slope of the line represents the scale factor value at each wavelength
Scale factors represent a percent change per a 1% change in Mg-II Index
Scale factors between the three OMPS NP sensors were expected to be similar, but not identical

## Slide 10
To determine the amount of correction that needs to be applied to each solar measurement, scale factors are multiplied by the percent change in Mg-II Index compared to the day-1 measurement
  For example, if the scale factor value is 0.5 and the change in Mg-II Index from day-1 is 4%, then the total impact from solar activity would be 2%
Wavelength shift adjustments are also made relative to day-1
  Shift adjustments are made to the data following adjustments for solar activity
Corrections have been calculated for the lifetime of each of the three NP sensors
Applying Both Corrections to the Data
Oldest Data  Most Recent Data
Oldest Data  Most Recent Data
Original
Solar Activity Corrected
SA and Shift Corrected
Original
Solar Activity Corrected
SA and Shift Corrected
(Animated for all Wavelength Channels)

## Slide 11
Calculating Sensor Optical Throughput Degradation
After correcting for both solar activity and wavelength shift in the working and reference diffuser measurements, degradation rates can be calculated
A curve fit is applied to the working and reference degradation rates
Working diffuser degradation rates are scaled based on the ratio between number of measurements from each diffuser
The difference between the reference diffuser degradation and the scaled working diffuser degradation is the throughput degradation
Goniometric discrepancies in reference diffuser measurements starting in 2022 were addressed by approximating the degradation between 2022 and 2025 based on the average yearly degradation between 2015 and 2021
Degradation adjustments are made relative to Reference Diffuser Measurement 3, as the first and second measurements are out-of-family with all following measurements
Working Diffuser
Reference Diffuser
Working Scaled to Reference
Throughput Degradation Estimate
Working Diffuser
Reference Diffuser
Working Scaled to Reference
Throughput Degradation

## Slide 12
With solar activity and throughput degradation known, the SNPP OMPS NP SDR solar flux data product can be adjusted to create a more accurate dataset
Wavelength shift adjustments are made every 2 weeks via LUT updates – solar activity and degradation adjustments will also be made at the same time
While degradation adjustments will be incorporated into bi-weekly LUT deliveries, the magnitude of the adjustment will remain constant until a new reference diffuser measurement is taken, usually once per year
Solar activity adjustments will vary with each new delivery
Adjusting the SNPP OMPS NP Solar Flux Data Product
Operational NPP NP SDR Solar, SDR Solar + SA Adjustment, SDR Solar + SA + Degradation

## Slide 13
Reprocessing efforts for SNPP OMPS NP SDR data are currently underway
Current comparisons of completed reprocessed data indicate that the applied adjustments have been successful
Reflectance values over the tropical Pacific region have been compared between operational and reprocessed datasets going back to 2019, with differences consistent with what was anticipated from the solar updates
Updated Normalized Radiance values have been projected from launch to present day by manually updating solar values. Operational and Reprocessed NR values were also compared over the tropical Pacific region. Results are also as expected
Improves the mission-long data quality of SNPP OMPS NP SDR Data, which better serves the GSICS community
SNPP OMPS NP SDR Data Reprocessing
Operational NR, Reprocessed NR

## Slide 14
Summary and Conclusions
The operational versions of the OMPS SDR solar flux data products only account for wavelength shift, and thus do not accurately represent changes in solar flux
Solar activity and optical throughput degradation impacts were determined for the SNPP OMPS NP sensor
This analysis provides a comprehensive update to the SNPP OMPS NP SDR solar flux data product, creating a more robust dataset that takes into account both solar activity and optical throughput degradation
These adjustments are shown to have an impact roughly between -4% and 6% on the solar flux data, varying both by wavelength channel and in time
Reprocessing of the full SNPP OMPS NP SDR data record is under way to update the solar flux values, with testing of the new data ongoing. This will be the third full reprocessing of the SNPP OMPS NP SDR dataset
Current analysis of reprocessed data has shown that impacts on Normalized Radiance and Reflectance values can also be up to roughly 6%, in line with what is seen in the solar flux differences

## Slide 15
Path Forward
Solar Flux updates will go into operations after the conclusion of the current reprocessing efforts and following coordination with the NOAA OMPS EDR team
  The NOAA OMPS EDR team currently makes soft calibration adjustments to the solar flux dataset, which will need to be adjusted to account for the updated SDR solar flux values
Work is ongoing to apply similar adjustments to NOAA-20 OMPS NP as well, with hopes of finalizing the adjustments soon
Plans are in place to extend the analysis to NOAA-21 OMPS NP, but only preliminary work has been done at this point
Work is also ongoing to calculate and analyze similar impacts on each of the three OMPS NM instruments, starting with SNPP OMPS NM, though the impacts are projected to be much smaller than those of NP

## Slide 16: Acknowledgements
Thanks to Mr. Eric Beach for collecting historical OMPS SDR data sets that are used in the study.
Thanks to the NASA OMPS Team (Glen Jaross, Tom Kelly, and Colin Seftor) for their development and contribution of RawFlux data and for their continued help in analysis of OMPS data

## Slide 17: Disclaimer
The presentation contents are solely the opinions of the authors and do not constitute a statement of policy, decision, or position on behalf of NOAA or the U. S. Government.

## Slide 18: Thank You!
Questions?

## Slide 19: Backup Slides

## Slide 20
NPP NP Tropical NR And Solar Flux Percent Differences

## Slide 21: Normalized Radiance and Reflectance PD from Reprocessed Dataset

## Slide 22
NPP NP Working and Reference Diffuser Correction Steps

## Slide 23: N20 NP Degradation Status