## Slide 1
Tropospheric Emissions:
Monitoring of Pollution
TEMPO Level 1 Status
Heesung Chong1, X. Liu1, D.E. Flittner2, J. Houck1, J. Carr3, W. Hou1, J. Park1, G. González Abad1, C.R. Nowlan1, H. Wang1, L. E. Flynn4, S. Marchenko5, J. Joiner5, J. E. Davis1, R. M. Suleiman1, J. Fitzmaurice1, L. Yatsuhashi1, K. Chance1, B. Baker6, C. Seftor5, G. Jaross5, J. O. Robaidek7, X. Wu4, H. Xu4, B. Basnet4, Y. Bai4, C. Cao4, D. Doelling2
1SAO, CfA 2NASA LaRC 3Carr Astro. 4NOAA/NESDIS 5NASA GSFC 6BAE 7UW–Madison
March 25, 2026

## Slide 2: Contents
Major updates
  Absolute calibration
  Stray light correction
  Spectral calibration
Additional updates
  Diffuser goniometry correction
  Etalon fringe correction
  Parity normalization
  Dark current correction
  Smear correction
Updates in Version 4
Major tasks
  Dynamic bad pixel detection
  Degradation correction
  Polarization correction
Additional tasks
  Spectral calibration update
  Stray light correction update
  Absolute calibration update
Toward Version 5

## Slide 3: Updates in Version 4

## Slide 4: Absolute calibration
Issue in Version 3: Positive bias in Sun-normalized radiance (on the order of 10%)
Approach: Adjust the solar diffuser transmittance.
(Earth radiance)
(Solar irradiance)
(Sun-normalized radiance) =
(Earth-radiance electric current) * (Calibration coefficient)
(Solar-irradiance electric current) * (Calibration coefficient) / (Diffuser transmittance)
=

## Slide 5: Absolute calibration
1024th spatial index
2023/08/01 (first-light)
(Solar irradiance) = (Solar-irradiance electric current) * (Calibration coefficient) / (Diffuser transmittance)
(Earth radiance) = (Earth-radiance electric current) * (Calibration coefficient)
(TEMPO Version 3 irradiance) – (TSIS-1 reference irradiance)
Why did the transmittance bias not appear in Version 3?  Biases in calibration coefficients
How to verify the hypothesis?  Earth radiance

## Slide 6: Absolute calibration
Applying in Version 4 has decreased Sun-normalized radiances by ~13.4%.
Solar diffuser
Definition of BTDF at GSFC during on-ground calibration
Definition of BTDF in TEMPO algorithm
Update in BTDF (Bidirectional Transmittance Distribution Function) data interpretation
Use of in TEMPO algorithm - Version 3: - Version 4:

## Slide 7: Absolute calibration
In Version 4, the UV and visible calibration coefficients have been reduced by 11% and 7%, respectively.
Earth radiance intercomparison (TEMPO V03 vs. ABI vs. VIIRS)
Example: 470 nm (UV), GOES-19

## Slide 8: Absolute calibration
Version 3
*Note: Spectral-independent, constant surface albedo (0.074) was used.
[Credit: Junsung Park]
2023/11/06, scan 009, granule 02
Sun-normalized radiance (Version 3  Version 4)
Version 4
Most significant outcome: Cloud fraction bias have been significantly reduced.
Ratio: TEMPO / VLIDORT simulation

## Slide 9: Stray light correction
Issues in Version 3:
Overcorrection at short wavelengths
Background stray light remaining in the UV channel
Undercorrected stray light in cloudy scenes
Most significant outcome: Ozone profile retrieval has been improved.
Improvement 1
Update PSFs (Point Spread Functions).
Improvement 2
Implement a background stray light correction.
Before
(308.03 nm)
After
(308.03 nm)
VIS PSF

## Slide 10: Stray light correction
Before reconstruction
After reconstruction
Improvement 3: Reconstruct radiances for CCD pixels with signal saturation to compute the source for the stray light.
Without reconstructing these flat-tops, stray light is underestimated.
Most significant outcome: Positive biases in formaldehyde retrieval over cloudy pixels have been mitigated.
(Reconstructed radiances are only for stray light correction and are not saved in Level-1 files.)
HCHO
fitting
window
HCHO
fitting
window

## Slide 11: Spectral calibration
Issue in Version 3: Spatial variability (North-South) in wavelength registration
Solution: Remove the slit-function asymmetry factor from the solar fitting to increase stability.
[Image credit: Gonzalo González Abad]
NO2 fitting window
HCHO fitting window
Spectral shift [nm]
Spectral shift [nm]
NO2 fitting window
HCHO fitting window
CCD spatial index (North ➝ South)
CCD spatial index (North ➝ South)
Version 3
Version 4
Most significant outcome: The Level 1 calibration capture spectral changes with enhanced accuracy.

## Slide 12: Toward Version 5

## Slide 13: Dynamic bad pixel detection
Current approach
Static bad pixel map
Quality flags for saturation, offset, nonlinearity, gain, and smear
Level 0-1 processor
Processing for dark current?
Quality flags for hot and cold pixels
Dark current
(50 frames)
Radiance
or
irradiance
No
Yes

## Slide 14: Dynamic bad pixel detection
New approach under consideration
Static bad pixel map
Dynamically updated bad pixel map
Quality flags for saturation, offset, nonlinearity, gain, and smear
Level 0-1 processor
Processing for dark current?
Quality flags for hot and cold pixels
Dark current
(50 frames)
Radiance
or
irradiance
No
Yes
if persistent
Most significant outcome: Level 2 retrievals are expected to improve by using only reliable CCD pixels.
Quality flags for transient

## Slide 15: Degradation correction
Solar irradiance trending (2023/08 – 2026/03)
UV (290–490 nm)
VIS (540–740 nm)
Working diffuser (every week)
Reference diffuser (every 3 months)

## Slide 16: Degradation correction
UV (290–490 nm)
=
+
Working irradiance change
Working diffuser change
Optics/detector change
To be corrected separately
Correction to be applied to every wavelength in the UV and VIS CCDs.
Most significant outcome: Level-2 retrievals using Sun-normalized radiances will be improved (O3 & cloud).
(estimated by working/reference irradiance ratio)
(estimated by reference irradiance)

## Slide 17: Polarization correction
Level-1
Sun-normalized radiance
Total ozone column
Estimate atmospheric and surface parameters
if LER:
Scene albedo & Scene pressure
if MLER:
Cloud fraction & Cloud pressure
Derive Stokes parameters
(Q & U)
Derive (1) degree of linear polarization and (2) phase angle of polarization
Polarization-corrected Level-1 spectrum
Earth radiance
Source of polarization: Atmosphere & Earth surface
Most significant outcome: Fitting residuals in Level 2 retrievals are expected to decrease.

## Slide 18: Polarization correction
Derive (1) degree of linear polarization
Polarization-corrected Level-1 spectrum
Solar irradiance
Source of polarization: Solar diffuser
Derive (2) phase angle of polarization
(two-plate fused silica)
Fresnel Equations
Incident elevation angle
Incident azimuth angle
Most significant outcome: Fitting residuals in Level 2 retrievals are expected to decrease.

## Slide 19: Summary
In Version 4, we enhanced
  Absolute calibration
  Stray light correction
  Spectral calibration
  Diffuser goniometry correction, etalon fringe correction, parity normalization, dark current correction, and smear correction.
For Version 5, we will enhance
  Bad pixel detection
  Degradation correction
  Polarization correction
  Spectral calibration, stray light correction, and absolute calibration

## Slide 20
Thank you for your attention!